Test operation of a novel Time-Of-Flight mass spectrometer on the gas exhaust of Wendelstein 7-X

G. Schlisio1, S. Gasc2, L. Hofer2, C.C. Klepper3, … and W7-X Teama

1 Max-Planck-Institut für Plasmaphysik, Wendelsteinstr. 1, 17489 Greifswald, Germany

2 Spacetek Technology AG, Brüggliweg 18, 3073 Gümligen, Switzerland

3 Oak Ridge National Laboratory, Oak Ridge, Tennesse 37831, USA

a for full team list see Pedersen NF2022, DOI…

Abstract

A novel time-of-flight mass spectrometer was operated in Wendelstein 7-X, a magnetic confinement fusion (MCF) experiment, to assess the suitability and limitations in the use for gas exhaust analysis in MCF devices. With a focus on high mass resolution sufficient for isotope separation, the permanent presence of magnetic field and a need for fast time resolution MCF presents a challenging environment for the operation of such devices.

…

Introduction

Mass spectrometers of various measurement principles are indispensable in many applications, vacuum quality monitoring and gas composition being a prominent one.

Wendelstein 7-X (W7X) is an optimized modular stellarator experiment dedicated to magnetic confinement fusion research in high-temperature plasmas [TODO reference].

Time-of-flight mass spectrometers (TOF-MS) measure mass to charge ratio m/z by accelerating ions against a retarding field and measure the response time in a detector. The flight time is directly proportional to the ions m/z ratio, which allows the mass-resolved measurement of the injected ions.

The IonTamer® FA (ITFA) is a novel TOF-MS <put properties here>. #TODO

Possible applications of a fast high-resolution mass spectrometer in magnetic confinement fusion (MCF) research and future fusion reactors:

- Gas exhaust monitoring, especially accounting of DT fuel

- He exhaust monitoring  divertor effectivity

- Impurity monitoring  assessment of plasma chemistry and vacuum quality

Residual magnetic field influence on the TOF-MS measurement

Many mass analyzers require free-streaming charged particles, and TOF-MS share this property. Fast free streaming particles are subject to a Lorentz force in a magnetic field. While MCF employs toroidal fields with a quadrupole far field characteristics, residual fields of several mT cannot entirely be avoided at reasonable distance of the torus system. Therefore, magnetic shielding is frequently employed to reduce the residual magnetic field down to an acceptable level.

Over the test duration, we successively improved the magnetic shielding in steps and document the effect below. While the residual field could not be measured in-situ, calculation of the residual field vectors were performed and the shielding structure was subjected to measurements in a test magnet setup thereafter.

Dynamic range dependence on pressure

The TOF-MS is designed for pressures up to 1e-3 Pa

Higher pressure means higher signal

Sudden changes in pressure can lead to safety deactivation of the instrument

(ANALYSIS HINT: take data from 20230307 where we reach transient saturation at some point and a short pumpdown at ~12:57CET reduces pressure and signal drastically; see also data taken on 2023-03-09: stepwise h2 pressure increase up to saturation, subsequent tuning-down of the gain, and pressure increase up to 1e-5mbar, gain scan at elevated pressures and a comparison with the ITMS and its spectra; see also 50ks=13.9h integration spectrum @~2e-8mbar 20230309)

Time response

Observing gas dynamics in the fusion exhaust requires a sub-second time resolution, which is often detrimental to dynamic range of instruments, as longer integration times give higher signal-to-noise ratio. The ITFA in its present

Observation of isotopes in common molecules

While W7X does currently only run Hydrogen, and neither Deuterium nor Tritium is present in the machine, the question of deuterated and tritiated molecules poses itself in view of Deuterium-Tritium fuel foreseen for reactors. W7X is planning for Deuterium operation in the next years, Tritium will be utilized in ITER and beyond. Identifying fuel-carrying molecules and precisely accounting for fuel will be a relevant and challenging part of the fusion fuel cycle.

While new developments [6] promise to extract over 90% of the Tritium via direct internal recycling, the remaining mixture remains to be analyzed and monitored.

Due to the unavailability of D and T in the experimental setup, simulations have been performed and are presented here along with measurements as a prediction of the TOF-MS capabilities in this regard.

Other notable findings

During operation of the ITFA in hydrogen-dominated sampling gas a prominent mass peak at m/z=3.024 was reliably observed. This corresponds to thrice the hydrogen mass and relates to H3+, a molecule long known and special relevance in astrophysics. According to literature, it forms with a proton transfer process H2 + H2+  H3+ + H which occurs in the ITFA analyzer. This effect is well-documented [4] but not common knowledge.

Discussion

Future development of the ITFA and its analysis tools are planned, with features like mass range selection.

Declaration of the authors

The authors declare that SG and LH are employed by Spacetek Technologies AG, manufacturer and marketer of the discussed ITFA.

Acknowledgements

The authors with to acknowledge the support of S. Vartanian and E. Gauthier for discussions on H3+. The authors further wish to acknowledge the fine lab work of A. Graband in support of the ITFA test, especially with regard to the magnetic shielding.

This work has been carried out within the framework of the EUROfusion Consortium, funded by the European Union via the Euratom Research and Training Programme (Grant Agreement No 101052200 — EUROfusion). Views and opinions expressed are however those of the author(s) only and do not necessarily reflect those of the European Union or the European Commission. Neither the European Union nor the European Commission can be held responsible for them.

References

[1] S. Gasc, L. Hofer, Chimia 2022, DOI: 10.2533/chimia.2022.52

[2] G. Schlisio et al, RSI 2019, DOI: 10.1063/1.5098125

[3] G. Schlisio et al, NF 2021, DOI: 10.1088/1741-4326/abd63f

[4] E. Gauthier, (1995). Quantitative analysis of deuterium and helium after plasma shots by means of resolved mass spectrometry (EUR-CEA-FC--1557). France (https://inis.iaea.org/search/search.aspx?orig_q=RN:27073670)

[5] McCall Phil.Trans.R.Soc.Lond.A 2000, DOI: 10.1098/rsta.2000.0655 (H3+ spectroscopy)

[6] Day FusEngDes2019 https://doi.org/10.1016/j.fusengdes.2019.04.019 or Haertl FusEngDes22 https://doi.org/10.1016/j.fusengdes.2021.112969